Evaluation of the Hycel Diagnostics Diana-5 Hematology Analyzer

Author(s): F. Boemer 1 ¦ N. Cielniazek 2 ¦ J. M. Paulus 3 ¦ A. Gothot 4

Laboratory Hematology

Print ISSN: 1080-2924 Volume: 8 ¦ Issue: 3 Cover date: 2002 Page(s): 126-133

Keywords. Diana-5 hematology analyzer, CBC parameters, WBC 5-part differential Abstract text. The Hycel Diagnostics Diana-5, a compact 5-part differential hematology analyzer, was evaluated in the hematology laboratory of Liège University Hospital, Belgium. Within and between-batch precision, linearity, and absence of carryover were all excellent. All parameters were stable for 48 hours at 4°C and for 8 hours at room temperature. Comparability of complete blood counts to Bayer Technicon Advia 120 measurements was highly satisfactory, with the exception of the mean cell volume, which was slightly under-estimated by the Diana-5 compared to the Advia 120. Comparisons of differential leukocyte counts to 400-cell microscopic differentials were excellent for neutrophils, lymphocytes, and eosinophils and were acceptable for monocytes. The efficiency of the flagging system was high for the detection of blasts and less satisfactory for the detection of immature granulocytes and atypical lymphocytes. The instrument was reliable and easy to use.

Author(s) affiliations

1. Department of Laboratory Medicine, Laboratory Hematology, University of Liège, Liège, Belgium 2. Department of Laboratory Medicine, Laboratory Hematology, University of Liège, Liège, Belgium 3. Department of Laboratory Medicine, Laboratory Hematology, University of Liège, Liège, Belgium 4. Department of Laboratory Medicine, Laboratory Hematology, University of Liège, Liège, Belgium

Traduzione dal testo:

Il sistema diagnostico Diana-5, un analizzatore differenziale compatto di Hycel di ematologia 5-part, è stato valutato nel laboratorio di ematologia dell'ospedale dell'università de Liegi, Belgio. La precisione, linearità ed assenza di carry over era tutta eccellente. Tutti i parametri erano stabili per 48 ore a 4°C e per 8 ore alla temperatura ambiente. La comparabilità completa dei conteggi con il Bayer Technicon Advia 120 misure era altamente soddisfacente, con l'eccezione del volume medio delle cellule, che è stato un po’ sottovalutato dal Diana-5 rispetto al Advia 120. I confronti dei conteggi differenziali dei leucociti ai differenziali microscopici 400-cell erano eccellenti per i neutrofili, i linfociti ed i eosinofili ed erano accettabili per i monociti. L'efficienza del sistema per la determinazione era alta per la rilevazione della lettura e soddisfacente per la rilevazione dei granulociti acerbi e dei linfociti atipici. Lo strumento era semplice e facile da usare